## Instructions
Checking the Temperature Calibration
July 21, 2009
1. Connect a digital voltmeter directly to the thermocouple (see Figure 1 and
2 below; NB! Be careful not to short circuit the thermocouple when the test pins
are connected)
2. Place a thermometer on the thermocouple connections to measure the
temperature of the connector block
3. Open the reader and remove the sample carrousel. Close the lid again
4. Open the \Control" program
5. Connect to MiniSys
6. Select the \Services" tab
7. Switch on the Nitrogen and lift the heater plate
8. In the eld \Set Temperature To" enter \50" and press \Go". The Ther-
mocouple will now ramp the temperature to 50C.
9. Read the voltage on the voltmeter and the connector temperature on the
thermometer
10. Convert the voltage reading (x) into temperature. This can be done by
using the Excel Spreadsheet \Temp calibration Check". Enter the connec-
tor temperature (cell E3), the set temperature (i.e. \50" in cell F6) and
the voltage reading (x) (in cell F7): The spreadsheet calculates the actual
temperature using a 6th degree polynomial with the following parameters:
x6 : -0.0000006782449
x5 : 0.00005292144
x4 : -0.001319532
x3 : 0.008238798
x2 : 0.06280513
x1 : 23.9038
x0 : 0.8888392
These parameters were determined using the ITS-90 table for type K ther-
mocouple.
11. Repeat the measurement at 100C, 200C, etc.
12. The spreadsheet \Temp calibration Check" plots the deviation of actual
temperature to set point temperature. The deviation should not be more
than approximately 1:0C. The deviation at 700C may be considerably
higher if the reader has been calibrated for 50 to 650C operating range.

## Calibration check
| Set temp (C) | 50 | 100 | 200 | 300 | 400 | 500 | 600 | 650 | 700
After linearisation | Reading (mV) | 0.99 | 3.05 | 7.11 | 11.17 | 15.32 | 19.54 | 23.82 | 25.94 | 27.93
Amb.Temp (C) | Reading (C) | 49.6219 | 99.5127 | 199.4829 | 299.556 | 399.6645 | 499.0512 | 599.3961 | 649.4929 | 696.544
25 | Deviation (C) | -0.37807 | -0.48734 | -0.51711 | -0.444 | -0.33546 | -0.94879 | -0.60387 | -0.5071 | -3.456